# Laser Capture Microdissection

- Use the Veritas machine, (or palm zeiss when available), with membrane slides, pretty much as per standard protocols, but with following considerations.

- Too strong a cutting laser might degrade RNA according to Chris Erwin. So cut back on power as much as possible. I’ve had good luck with the setting at 3.0.

- Be careful to avoid cap crap. But if there is a little then can use ablation to remove. But be careful to re-set laser after ablation, so don’t ruin sample.

- Label caps clearly and take photos before, after and of caps to record material captured.

- Place caps in 0.5 ul tubes after capture is complete.

- Store on dry ice, then at –80ºC.